## REM SYSTEMS PTY LTD

## NIKI T34 Setup

Please follow the  instructions below when  using the NIKI T34 syringe driver for fixed 24 hour infusions. The following instructions are relevant where the NIKI T34 has been pre-programmed to a fixed duration of 24 hours and has the Program Lock function on.

DAY 1

- Prepare medication in relevant syringe and attach infusion line and catheter – do not attach infusion line to the patient, yet

- Turn on NIKI T34

- Measure syringe against NIKI T34 and press either  FF  or  Back  to align actuator to syringe plunger.  Actuator will re-set to previous syringe size so if same size syringe no need to alter

- Raise Clamp and place syringe in, close clamp, check 3 syringe sensors.

- Use  to Select syringe  brand.

- Press NO for new patient.

- Check and review data on screen, Volume, Duration, Rate, (eg 24ml, 24hrs, 1ml/hr)

- To confirm Press YES

- Display will read “ START INFUSION”.DO NOT  commence infusion.

- Press the FF key. The display will read “PURGE. DISCONNECT FROM PATIENT”

- Confirm by pressing  YESC

- Prime Infusion line and catheter  by pressing and holding the key FF  ( up to 2ml)

- Prime is now complete.

- Select correct brand of syringe.

- Press  YES to Resume.

- Display will now read  new  volume (eg22.5ml) and new Duration (eg. 22.5hrs) Rate will remain constant (eg 1ml/hr)

- To Confirm press  YES

- Connect infusion line to patient

- When “Start Infusion?” is displayed, to commence  press  YES

DAY 2, 3 etc

- Prepare medication in relevant syringe

- Turn off NIKI T34 pump

- Remove old syringe and attach new syringe to infusion line.

- Measure syringe against NIKI T34 and press either  FF  or  Back to align actuator to syringe plunger

- Raise Clamp and place in syringe, close clamp

- Select correct syringe brand

- Check data on screen, Volume, Duration, Rate, (eg 24ml, 24hrs, 1ml/hr)

- To confirm Press  YES

- When “Start Infusion?” is displayed, to commence Press YES